# Usability study variables

| Variable / Information | Details | Level | Time | Source |
| --- | --- | --- | --- | --- |
| Inclusion/exclusion criteria |  | Professional CT | Pre-selected | eCRF |
| Characteristics of professional | Age range, number interventions/year | Professional | Baseline | eCRF |
| SUS Scale (10 items) | 10 Likert Items, and summary score. | Professional | Post- review | eCRF |
| SUS specificity | 8 Likert items | Professional | Post- review | eCRF |
| Qualitative Feedback | 4 open questions | Professional | Post- review | eCRF |
| Time analysis interaction | The overall workflow - Load a new case -Perform measurements (CT) - Select the landing zone - Choose device size - Execute a device deployment -Review the final report | Professional CT (Case) | Per case review | eCRF Platform – eCRF Platform – eCRF Platform – eCRF Platform – eCRF Platform – eCRF Platform – eCRF |
| DRT risk stratification | Binary option (Yes/No). Patient-specific hemodynamic simulations based risk classification improved my understanding of DRT patient risk | Professional – CT (Case) | Per case review | eCRF |
| Decisional change | Binary option (Yes/ No). Implant chosen differs from that used in standard-of-care. | Professional -CT (Case) | Per case review | eCRF |
| Critical Task* | 8-item checklist covering platform task execution and reporting questions | Professional -CT (Case) | Per case review | eCRF |

Critical Task extretes de taula 5 (protcol)

Cada tasca prevista  segons escenaris.

En el context de la plataforma per a la planificació prèvia del tancament percutani de LAAO i l’estratificació del risc de DRT, un error crític d’ús és aquell que probablement:

• Donaria lloc a la selecció d’una mida de dispositiu inadequada, amb el risc d’un tancament ineficaç;

• Comportaria una interpretació errònia dels resultats de la simulació del tancament del LAA, conduint a decisions clínicament perilloses;

• Implicaria que l’usuari no identifiqués un pacient amb un perfil d’alt risc de DRT

Nota en taula 4 hi ha un llistat més ampli (13) però inclou tasques no critiques I son més de 8.

S’enten que s’ha de valorar si cada tasca es completa correctament

| Use Scenario | Critical Task | Associated Use-Related Risk | Potential Hazard | Mitigation Measures |
| --- | --- | --- | --- | --- |
| 1. Pre-procedural planning of LAAO | 1. Interpret simulation outputs (wall apposition maps, device deformation). | Misinterpretation of sealing performance or device fit. | Incomplete LAA closure. | Tutorial overlays, tooltips, training module, task walk-through, IFU. |
|  | 2. Select device size and position. | Selection of wrong device size and position. | Undersizing. Oversizing. Incomplete LAA closure | Visual alerts guiding optimal positioning, recommendations based on sizing guides and compression. |
|  | 3. Identify predicted peri-device leak. | Missed leak prediction due to oversight or misreading. | Residual flow that leads to thrombus formation. | Binary color codes detecting incomplete sealing areas leading to possible leak. |
| 2. Post-implant DRT risk evaluation | 4. Interpret patient-specific hemodynamic simulations output (ECAP), average velocity and particle attachment. | Misclassifying patient DRT risk | Incorrect clinical decision | ECAP threshold alerts, visual annotations, training module, on-screen description of parameters and results. |
|  | 5. Assign DRT risk category (high vs. low) | Inaccurate or reversed risk assignment. | Incorrect clinical decision | Binary classification Summary. |
| 3. Interaction with software platform | 6. Navigate between modules. | User cannot locate relevant data. | Delayed or incomplete decision-making. | Training and IFU guidance. |
|  | 7. Complete decision-making within allotted time. | Delays due to confusion or poor interface. | Time pressure leading to cognitive overload. | Support cues, minimal UI design. |
| 4. Clinical comparison to SoC | 8. Provide rationale for changed decisions. | Failure to verbalize or document clinical judgment. | Loss of audit trail or unclear benefit. | Structured feedback form. |

# Concordance study

S’utilitzaran dades clíniques de la historial clínica dels pacients que prèviament van sotmetre’s a LAAO amb un dels dos dispositius utilitzats. Part de la informació del CRD provindrà d’una font secundaria.

L’elegibilitat dels pacients es basa en la disponibilitat de TC pre- i postprocedimentals d’alta qualitat, aptes per a reconstrucció 3D, així com dades Doppler/TTE 2D amb forma d’ona pulsada adequades per a l’anàlisi de simulacions hemodinàmiques específiques per pacient.

Concordance study variables

| Information/ Variable | Details | Time | Source |
| --- | --- | --- | --- |
| Inclusion/exclusion criteria | Age>=18 | Pre selected | EHR - eCRF |
|  | Implanted device (o Watchman FLX / Amplatzer Amulet ) | Pre selected | EHR - eCRF |
|  | Availability of high-resolution CT scans | Pre selected | EHR - eCRF |
|  | Availability of procedural metadata | Pre selected | EHR - eCRF |
|  | Missing documentation of implant type, device size | Pre selected | EHR - eCRF |
| Anatomical endpoint Device geometry | - Frame cross-sectional area (mm) simulation - Frame cross-sectional area (mm) real | Post CT load Post review CT /EHR | VIDAA prediction CT pre-post CT Post |
|  | -Maximum diameter simulation -Maximum diameter real | Post CT load Post review CT /EHR | VIDAA prediction CT- EHR |
|  | -Minimum diameter simulation -Minimum diameter | Post CT load Post review CT /EHR | VIDAA prediction CT - EHR |
|  | -Perimeter simulation -Perimeter real | Post CT load Post review CT /EHR | VIDAA prediction CT - EHR |
| Leak endpoint | Peri-device leak prediction (Yes / No) Peri-device leak event (Yes/No) | Post CT load Post review CT /EHR | VIDAA – CRF CT- EHR |
| DRT | Patient-specific hemodynamic simulations based risk (low / high)* DRT event (Yes/No) | Post TC load Post review CT /EHR | Platform – CRF EHR |
| Basal characteristics of patients | Age, sex, BMI | Baseline | EHR -CRF |
| Clinical risk scores | CHAD2DS2-VASc score HAS-Bled score | Baseline Baseline | EHR -CRF |
| AF type | Paroxysmal / Permanent | Baseline | EHR -CRF |
| History of the patient | Heart failure (Yes / No) Diabetes (Yes / No) (Type 1) Ischemic stroke or TIA Other vascular disease Hypertension Abnormal liver function Abnormal renal function Ischemic heart disease PCI CABG LVEF (5) Valvular heart disease CRT/pacemaker/ICD | Baseline | EHR -CRF |
| Indications | Intracranial hemorrhage GI bleed Urinary tract bleeding, Other spontaneous bleeding Stroke despite OAC Cerebral amyloid angiopathy Cognitive impairment Patient preference/lack of compliance/side effects | Baseline | EHR-CRF |
| Device related data | Type (Watchman FLX, Amplatzer Amulet). Size Procedural implantation details (e.g., approach, depth, deployment characteristics). | Post intervention | EHR-CRF |
| Treatment | Discharge anti-thrombotic regimen (Yes/No) | Post intervention | EHR-CRF |
| CT Segmentation Validation | Subset of CTs (15-20) -> Manual segmentation by biomedical engineer → Clinical review (OK / Modify) → Comparison with automatic segmentation: DSC, Hausdorff distance 95. | Post CT load | Platform – CRF |

Table 8 – Concordance study variables